## Slide 1: Pre-processing methods for spectral data

## Slide 2: Why pre-process?
NIR spectral data are almost always pre-processed. The motivation for this is that the spectra tend to be affected by scattering or pathlength effects that contribute more variability to the spectra than the chemical composition we are usually trying to measure. Calibration methods like PCR and PLS find it much easier to identify the important factors if we first remove some of these effects.

## Slide 3: An example
This plot shows 49 NIR transmission spectra of wheat measured in a Tecator instrument. Most of the variability in the spectra reflects physical rather than chemical properties of the wheat.

## Slide 4
Loadings from a PCA of the 49 raw spectra.
Blue PC1 98.96%
Green PC2 0.99%
Red PC3 0.03%
The first two come from variability in the level and slope of the spectra. The third may be related to the composition of the wheat.

## Slide 5
Cross validation for the number of PLS factors for a protein calibration

## Slide 6
PLS calibration with 6 factors, RMSECV=0.35

## Slide 7
The 49 wheat spectra after pre-treatment with second derivative and then SNV. Now we can see some compositional information in the spectra. We can also see an outlier.

## Slide 8
Loadings from a PCA of the 49 2nd derivative SNV-treated spectra
Blue PC1 85.61%
Green PC2 4.81%
Red PC3 2.28%
These are not only more interesting, but relate to composition.

## Slide 9
Cross validation for number of PLS factors for a protein calibration.

## Slide 10
PLS calibration with 2 factors, RMSECV=0.28

## Slide 11: Common pre-treatments
Derivatives - order: first, second, ...
- method: Savitzky-Golay, Norris
Multiplicative - SNV, MSC

## Slide 12: First and second derivatives
The first of the 49 spectra (blue), with its first (green) and second (red) derivatives

## Slide 13: Need for smoothing
Raw spectrum and second derivative calculated by differencing twice

## Slide 14: Savitzky-Golay smoothing and derivatives
The basic idea is to fit a low order polynomial (quadratic or cubic) to a moving window of data points and use its value (to smooth) or its derivatives at the middle point.
In the case of a quadratic, the value at wavelength w is
y = a + b.w + c.w2
the first derivative is
dy/dw = b + 2c.w
and the second derivative is
d2y/dw2 = 2c

## Slide 15
The 100 individual data points of one of the wheat spectra with a 7-point window in red.

## Slide 16
Quadratic fitted to the 7 data points from 936 to 948nm. red dotted line is slope at 942nm.

## Slide 17: The effect of window width
Second derivative using windows of width 3, 5, 7, 9 and 25. Clearly 25 is too wide. The next slide expands the region around 1012nm.

## Slide 18
In the trade off between smoothness and loss of detail, I’d go for a width of 5 (red)

## Slide 19: Norris derivatives
Sometimes called Norris-Williams derivatives or segment-gap derivatives
The idea is to apply a boxcar smooth and then calculate the derivative by differencing the smoothed values (repeatedly for derivatives higher than the first) across a specified gap. One needs to choose both the window width for smoothing and the gap. In some implementations a further smoothing may then be applied.

## Slide 20: Comparison of methods
Red is S-G, 7 point window, blue is Norris segment=3, gap=3
When the overall window widths are similar, the results are similar.

## Slide 21: Some comments on derivatives
Both of these methods require a fudge at the end if they are not to lose a few points.
I don’t think it matters much which method you use.
It isn’t obvious that the derivative that looks nicest will give the best calibration. I choose by looking at the derivative spectra and picking the narrowest window that doesn’t give a very rough looking result. Others (over?) optimise calibration performance. I would usually try both first and second derivative.

## Slide 22: Multiplicative pre-treatments
Derivatives remove additive baseline effects but NOT multiplicative effects. These are the first derivative spectra of the wheat samples.

## Slide 23: Multiplicative scatter correction (MSC)
Plot of wheat spectrum 1 versus mean of 49 spectra. Fitted line has equation
x = 1.10xmean  0.34
correction is
x -> (x+0.34)/1.10

## Slide 24
Raw spectra
MSC-treated spectra

## Slide 25: Standard Normal Variate (SNV)
Raw spectra
Means subtracted
SNV standardises each spectrum to have mean 0 and standard deviation 1

## Slide 26
Means subtracted
SNV-treated

## Slide 27: MSC and SNV usually give similar results
SNV-treated
MSC-treated
They are however NOT the same. See, eg, Chemolab 96, 22-26, 2009.

## Slide 28: More than one pre-treatment?
It is possible, and may be beneficial, to apply more than one pre-treatment. In particular, a multiplicative scaling is often used together with a derivative.
It is not generally appreciated that the order of application matters. Nor, amongst those who do realise it matters, is there any general agreement about which is the ‘correct’ order.

## Slide 29: The order matters
D2 then SNV
SNV then D2

## Slide 30: Effect on calibrations?
Treatment | # PLS factors | RMSECV
Raw | 6 | 0.35
SNV | 5 | 0.34
MSC | 4 | 0.33
D1 (5pt S-G) | 4 | 0.36
D1+SNV | 3 | 0.33
SNV+D1 | 4 | 0.33
D2 (5pt S-G) | 3 | 0.32
D2+SNV | 2 | 0.28
SNV+D2 | 3 | 0.31

## Slide 31: More on multiplicative treatments
Before SNV
After SNV
Pre-treatment may destroy useful information

## Slide 32
Before SNV
After SNV
Pre-treatment may shift information
The original MSC paper suggested a weighted regression, with known measurement peaks getting zero weight, but this is rarely done.

## Slide 33: Choosing a pre-treatment
I don’t have any easy answers to this very difficult problem, just some comments that may or may not be useful.
Most NIR calibrations benefit from some sort of pre-treatment.
For quantitative calibrations my impression is that there often isn’t much to choose between several options. My favourite is second derivative then SNV, with the narrowest window that gives a non-noisy derivative spectrum. Others will have their own favourites, and most of the time it won’t matter much.

## Slide 34
For any calibration method that uses distances between spectra, the choice of pre-treatment will often be crucial, and more effort needs to be put in to trying several. The most obvious example would be KNN for qualitative calibrations.
There exists software that will systematically search a large family of pre-treatments to find the optimal one for a particular data set.
There are two traps to avoid here:
Be wary of over-interpreting very small differences in performance and thus choosing an unnecessarily complicated pre-treatment
Don’t use the ‘independent validation set’ for this optimisation. If you do, its role as a validation set is compromised.

## Slide 35: A final comment
Pre-treatments do not add information, they remove it. They work by exposing the useful information underneath what has been removed. It may be helpful to think of the process as similar to that of cleaning a very dirty painting. Be careful, or you might spoil the picture.